## Slide 1: Spectrophotometry

## Slide 2: Spectrophotometry
It is that technique that measures the amount of light absorbed or transmitted by a substance.
It is one of the most important technique in analytical biochemistry.
Unknown compounds can be identified by their characteristic absorption spectra in the ultraviolet, visible or infrared regions.
Concentrations of known compounds in solutions can be determined by measuring the light absorption at one or more
wavelengths.

## Slide 3: Spectrophotometry cont’d
Spectrophotometry is used for both and quantitative and qualitative analysis .
Enzyme catalyzed reactions can be followed by measuring the absorption of the substrate or product .

## Slide 4: Spectrophotometry cont’d
The ultra violet and the visible regions are the ones that we usually use in the spectrophotometry.
1 nm = 10-9 m
1 A° = 10-10 m
1 μm = 10-6 m
λ is the symbol of wavelength.
Regions | X-Ray | Ultraviolet | Visible | Infrared | microwave
Wave length | 0.1 -100 nm | 100 - 400 nm | 400 - 800 nm | 800 nm - 100 μm | 100 μm - 30 cm
Notes: Angstrom
Angstrom (Å), unit of length used chiefly in measuring wavelengths of light, equal to 10−10 metre, or 0.1 nanometer.

## Slide 5: Spectrophotometry cont’d
Wave number is the reciprocal of the wavelength
wave number (cm-1) =
1
Wavelength

## Slide 6: The essential components of spectrophotometer
1- Light source:
  It can be two kinds:
    Tungsten lamp ; produces light at visible region.
    Hydrogen lamp; produces light at ultraviolet region.
2- Collimator:
  It is a focusing device that transmits an intense straight beam of light.
3- Monochromator:
  It is a device that divides the light beam into it’s component wavelengths.

## Slide 7: The essential components of spectrophotometer cont’d
4- Selector:
  It selects the required wavelength.
5- Cuvette:
  It is a compartment in which the sample is placed.
  Two kinds:
    Glass cuvettes; used in the visible region.
    Quartz cuvettes; used in the ultraviolet region.
    The glass cuvettes absorbs light in the ultraviolet region .. Thus the amount of light measured by spectrophotometer will be the absorbance of sample + the glass cuvette.

## Slide 8
6- Photocell (photodetector):
  It detects the amount of light transmitted.
7- Electrical meter:
  It records the output of the detector.
The essential components of spectrophotometer cont’d

## Slide 9
Light source
Monochromator
Prism
Light’s band
Cuvette sample container
Photocell
Slit
Wavelength selector
Potometer
Collimator
λ1
λ2
λ3
λ4
The essential components of spectrophotometer cont’d

## Slide 10
The essential components of spectrophotometer cont’d

## Slide 11: The essential components of spectrophotometer cont’d
The fraction of the incident light I0 that is absorbed by a solution depends on three factors:
  1- The thickness of the sample or path length.
  2- The concentration of the absorbing sample.
  3- The chemical nature of the compound.
The relationship between the concentration, path length, and the amount of light absorbed or transmitted can be exposed mathematically in two laws:
1-the Beer law and
2- Lambert law.

## Slide 12: Beer law
It states that the intensity of the light transmitted by an absorbing media decreases with increasing concentration of the absorbing compound.
Log I0 / I α a c
I0 = intensity of the incident light
I = intensity of the transmitted light
a = a constant for every compound at a specific wavelength
c = concentration of absorbing compound
Notes: the Beer-Lambert law states that there is a linear relationship between the concentration and the absorbance of the solution, which enables the concentration of a solution to be calculated by measuring its absorbance extinction coefficient or molar absorptivity

## Slide 13: Beer law cont’d
The amount of light absorbed = I0 – I
Log I0 / I represents the fraction of light absorbed

## Slide 14: Lambert’s law
It states that the intensity of the light transmitted by an absorbing media decreases as the thickness or path length of the absorbing media increases.
Log I0 / I α a L
L = is the path length
L

## Slide 15: Lambert-Beer law
Log I0 / I = a c L
a = is the extinction coefficient which is a constant for each substance but it differs at different wavelength and it is constant at a specific wavelength
  Molar absorption coefficient (am); in M
  Specific absorption coefficient (as); in g/l
  am340 = is the molar absorption coefficient of a substance at a wavelength = 340 nm
  am = as Mwt
am is most commonly used in biochemistry, and the path length L is almost always 1 cm, thus the units for am is M-1 cm-1.
The absorption coefficient varies in different substances, it also varies with varying wave-lengths also.

## Slide 16: Lambert-Beer law
Log I0 / I = A
  A = absorbance or optical density O.D.
  A = a c l

## Slide 17
The two law’s can be combined in one law which is Lambert-Beer Law
A = am x C x L
am= Molar absorption coefficient

## Slide 18: Blank solution
• Is a solution that is necessary in all spectrophotometry studies.
• It should contain all components of the assay or test solution except the component who’s absorbance is being measured.
Purpose of the Blank:
The blank will cancel out the absorbance of the substances in the background so that the absorbance of the
tests will be that of the compound under study only.
➢ Note: Glass cuvettes are not to be used in the U.V region, since the glass itself will absorb light thus leading to a false high result.
➢ In the U.V region Quartz cuvettes are to be used!

## Slide 19: Protein determinations
Proteins in solutions can be determined spectrophotometrically by two methods:
a) Colorimetrical method:
Biuret method: it is based on the reaction of Cu2+ with peptides in an alkaline solution producing a purple complex that has an absorption maximum at 540 nm.
Proteins + Biuret reagent purple complex (max absorbance at 540nm)
b) Direct spectrophotomety:
The absorbance at 280nm can be used to determine protein concentration in solutions.
(Because proteins have a distinct absorbance maximum at 280nm due to their aromatic amino acids).
alkaline media

## Slide 20: Coupled Assay
Many compounds of biological importance do not have a distinct absorption maximum λmax .
Their concentration can be determined if they can be linked to or coupled with a reaction that fulfills the following condition ..
If the promote the formation of another compound that has a characteristic absorption peak.
In coupled assay reactions the product of the first reaction is the substrate of the following reaction ;
A + B ------------> C + D ------------> E + F
A = substance under study that does not have a distinct λmax
F = has a distinct λmax
Thus A can be estimated by measuring the absorbance of F

## Slide 21: Coupled Assay Example
To 2.0 ml of a glucose solution; 1.0 ml of a solution containing excess ATP, NADP+, MgCl2, hexokinase and Glucose -6-phosphate dehydrogenase was added. Calculate the concentration of glucose in the original solution. Absorbance of final solution at 340 nm increased to 0.91, am= 6220, L = 1 cm.
Glucose + ATP. glucose -6-phosphate + ADP
Glucose-6-phosphate 6-phosphogluconate
+ NADP+ + NADPH + H+

## Slide 22: Coupled Assay Example
Glucose has no absorption at 340 nm, but NADPH does!
From the reaction: 1 mole of glucose produces 1 mole of NADPH, thus each number of NADPH moles produced originates from every mole of glucose in the original solution.
Absorbance at 340 nm is the absorbance of NADPH = 0.91
A = am x C x 1 = 0.91 = 6220 x C x 1
C NADPH = 0.91 / 6220 = 1.46x10-4 M
Thus there is 1.46 x 10-4 M of glucose present in the test solution
The glucose concentration in the original solution = 1.46 x 10-4 x (3/2) = 2.2 x 10-4 M 
Dilution factor

## Slide 23: Problems
Calculate the absorbance and the transmission at 260 nm and 340 nm of the following solutions in a 1 cm cuvette.
2.2 x10-5 M NADH
7 x 10-6 M NADH plus 4.2 x 10-5 M ATP.
This solution contains one absorbing substance (NADH)
A260 = am x C x l
A260 = 15000 x (2.2 x10-5 ) x 1 = 0.33
A = Log I° / I
0.33 = log 1.0 – log I
0.33 = - log I
I = antilog – 0.33 = 0.464
Absorbance and transmission at 340nm
A = 6220 x 2.2 x10-5 x1 = 0.1368
A = Log I° / I
0.1368 = log 1.0 – log I
am |  | 
Compound | 260nm | 340nm
NADH | 15000 | 6220
ATP | 15400 | 0.0

## Slide 24: Problems
0.1368 = -log I
I = antilog -0.1368
I = 0.729
b) The solution contains two absorbing substances
At 260nm
A = ANADH + A ATP
ANADH = 15000 x (7 x 10-6 ) x1 = 0.105
A ATP = 15400 x (4.2 x 10-5 ) x 1 = 0.646
A Total = 0.105 + 0.646 = 0.751
A= Log I° / I = 0.751 = log 1.0 –log I
0.751 = -log I , I = antilog - 0.751 , I = 0.177 .
At 340 nm only NADH absorbs
A = 6220 x (7 x 10-6 ) x1 = 0.043
A= Log I° / I = 0.043 = log 1.0 –log I , so I = antilog - 0. 043 , I = 0.905

## Slide 25: Problems
Calculate the concentration of ATP and NADPH in solutions with absorbance’s:
a) 0.15 at 340 nm and 0.9 at 260 nm.
b) Zero at 340 nm and 0.750 at 260 nm.
c) 0.22 at 340 nm and 0.531 at 260 nm.
a) Since this solution contains two absorbing substances , thus we will start with absorbance at 340nm since only NADPH absorbs
A 340nm = A NADPH only
A = am x C x l = 6220 x C x 1
C = 0.15 / 6220 = 2.4 x 10-5 M
am |  | 
Compound | 260nm | 340nm
NADPH | 15000 | 6220
ATP | 15400 | 0.0

## Slide 26: Problems
At 260nm:
A260nm = A ATP + A NADPH
A NADPH = am x C x l = 15000 x 2.4 x 10-5 x1 = 0.36
A ATP = A Total - A NADPH = 0.9 – 0.36 = 0.54
A ATP = am x C x l = 0.54 = 15400 x C x1
C = 0.54 / 15400 = 3.5 x 10-5 M
b) Since Absorbance at 340 nm is zero , and NADPH is the only absorbing substance at that wavelength thus the concentration of NADPH is zero
Accordingly the absorbance 0.75 at 260 nm is the absorbance of ATP only.
A ATP = am x C x l
C ATP = 0.751 / 15400 = 4.8x10-5 M
c) At 340nm only NADPH absorbs
0.22 = 6220 x C x1
CNADPH = 0.22/ 6220 = 3.5 x10-5 M

## Slide 27: Problems
At 260nm both ATP and NADPH absorb
Thus A = A ATP + A NADPH
A NADPH = 15000 x ( 3.5 x 10-5 ) x1 = 0.53
Since A NADPH = Atotal
Thus ATP concentration must be Zero

## Slide 28
A solution containing 2 g/l of a light absorbing substance in a 1 cm cuvette transmits 75% of the incident light at 260 nm.
Calculate the transmission of a solution containing
a) 4 g/l.
b) 6 g/l.
c) If the Mwt is 250, calculate am.
d) What type of cuvette should you use here? Why?

## Slide 29: Solutions Containing One Absorbing Substance
A = Log I0 / I
A = log 1.0 /0.75 = 0.124
A = as c l
as = A / c l = 0.124 / 2 × 1
as = 0.0625
A) log I° / I = as c l
Log 1.0 – log I = 0.06 x c x l
0 – log I = 0.0625 x c x l
- log I = 0.0625 x 4 x 1 = - 0.25
I = antilog - 0.25 = 0.562  56.2%

## Slide 30: Solutions Containing One Absorbing Substance
B) Log I0 / l = as c l
Log 1.0 - log I = 0.0625 x 6 x1.
- log I = 0.375
Log I = - 0.375
I = antilog - 0.375 = 0.422  42.2%
C) am = as x Mwt = 0.0625 x 250 = 15.63
D) Quartz cuvettes should be used at the U.V range.

## Slide 31: REVISION

## Slide 32: Solutions Containing One Absorbing Substance
A solution containing 10-5 M ATP, has a transmission 0.702 (70.2%) at 260 nm in a 1 cm cuvette. Calculate:
The transmission of the solution in a 3 cm cuvette.
The absorbance of the solution in a 1 cm and 3 cm cuvette.
The absorbance if the concentration increased to 5x 10-5 M of ATP, in a 1 cm cuvette.

## Slide 33: Solutions Containing One Absorbing Substance
a) A = Log I° / I = am c l
A = log 1.0 / 0.702 = 0.152
0.152 = am x 10-5 x 1
am = 0.152 / 10-5 = 15200 M-1 cm-1
A = 15200 x 10-5 x 3 = 0.456
A = Log I° / I
0.456 = log 1.0 / I
0.456 = log1 – log I = 0 – log I = - log I
Thus I = antilog - 0.456 = 0.349  34.9%
b) A in a 1 cm cuvette.
A = 15200 x 10-5 x 1 = 0.15
A in a 3 cm cuvette.
A = 15200 x 10-5 x 3 = 0.456
c) A = 15200 x (5 x 10-5) x 1 = 0.76

## Slide 34: Solutions Containing One Absorbing Substance
A protein solution (0.3 ml) was diluted with 0.9 ml of water. To 0.5 ml of this diluted solution, 4.5 ml of biuret reagent was added and the color was allowed to develop. The absorbance of the mixture at 540 nm was 0.18 in a 1 cm diameter tube.
A standard solution (0.5 ml containing 4 mg of protein/ml) plus 4.5 ml of biuret reagent gave an absorbance of 0.12 in the same size test tube.
Calculate the protein concentration in the undiluted unknown solution.
What is the composition of the blank here ?

## Slide 35: Solutions Containing One Absorbing Substance
A) Concentration of standard Cst = 4 mg/ml
Thus Cst = 4 g/L
Astandard = as x C x l
0.12 = as x 4 x 1
So as = 0.12 / 4 = 0.03
Atest = as x C x l
0.18 = 0.03 x C x1
So Ctest = 0.18 / 0.03 = 6 g/L = 6 mg/ml
The concentration of protein in the undiluted solution:
Cundiluted = 6 x (1.2/0.3) = 24 mg/ml .

b) The blank should contain 4.5 ml of biuret and 0.5 ml of distilled water only.
Dilution factor

## Slide 36: Solutions Containing Two Absorbing Substance
A solution containing NAD+ and NADH had an absorbance of 0.311 in a 1 cm cuvette at 340 nm, and 1.2 at 260 nm. Calculate the concentration of the oxidized and reduced forms of the coenzyme in the solution. Both NAD+ and NADH absorb at 260nm, but only NADH absorbs at 340nm.
Absorbance at 340nm represents the absorbance of NADH only since NAD+ does not absorb at that wavelength. So the concentration of NADH can be obtained.
A340nm = ANADH = am x C x l
0.311 = 6220 x C x 1
So CNADH = 0.311/6220 = 5 x 10-5 M
Absorbance at 260nm
A260nm = ANADH + ANAD+ ( since both absorb at this wavelength )
am |  | 
Compound | 260nm | 340nm
NAD+ | 18000 | 0.0
NADH | 15000 | 6220

## Slide 37: Solutions Containing Two Absorbing Substance
ANADH = am x C x l = 15000 x 5 x 10-5 x 1 = 0.75
Thus ANAD+ = A total - ANADH = 1.2 – 0.75 = 0.45
Since ANAD+ = am x CNAD+ x l
0.45 = 18000 x CNAD+ x 1
CNAD+ = 0.45 / 18000 = 2.5 x 10-5 M

## Slide 38
Below is a table prepared by a biochemistry student to construct a standard curve for protein analysis. The Bradford assay was used with bovine serum albumin (BSA, 0.1mg/ml ), as standard protein . Complete the table by filling in the approximate A595 that will be obtained for each tube. Assume the procedure was conducted correctly.
Reagents | 1 | 2 | 3 | 4 | 5 | 6
H2O (ml) | 1.0 | 0.9 | 0.8 | 0.6 | 0.2 | ----
BSA volume (ml) | ---- | 0.1 | 0.2 | 0.4 | 0.8 | 1.0
Bradford reagent (ml) | 5.0 | 5.0 | 5.0 | 5.0 | 5.0 | 5.0
A595 |  | 0.08 | ? | ? | ? | ?
Known BSA volume ---------------------Known Absorbance
0.1 ----------------0.08
0.2 ---------------- ?
Absorbance of tube 3 = (0.2 x 0.08) /0.1 = 0.16
(0.16)
(0.32)
(0.6)
(0.8)

## Slide 39: REVISION

## Slide 40
What is the net charge of the following amino acids:
a) Glutamate at pH 10 pKa1= 2.19 , pKa2 = 4.25 , and pKa3=9.67
b) Lysine at pH 1.5 pKa1= 2.18 , pKa2 = 8.95 , and pKa3=10.53
c) Serine at pH 5.8 pKa1= 2.21 , pKa2 = 9.15
d) Isoleucine at pH 1.0 pKa1= 2.36 , pKa2 = 9.68

## Slide 41
Choose the correct answer:
The amino acid arginine contains a guanidino R-group and has pKa values of 2.2, 9.0, and 12.5.
A sample of arginine is titrated from pH=1.0 to pH=14.0 with NaOH.
1._____ At pH=2.2,
a) all of the amino acid molecules will be in the fully protonated form.
b) half of the amino acid molecules will be in the fully protonated form.
c) all of the amino acid molecules will be in the zwitterion form.
d) half of the amino acid molecules will be in the zwitterion form.

## Slide 42
Cont’d
2._____ At pH=12.5,
a) half the amino acid molecules have a -2 charge.
b) all the amino acid molecules have a -2 charge.
c) half the amino acid molecules have a -1 charge.
d) all the amino acid molecules have a -1 charge.
3._____ For arginine molecules at pH=14,
a) all the ionizable groups will be charged.
b) all the guanidino groups will be charged.
c) all the amino groups will be charged.
d) all the carboxyl groups will be charged.

## Slide 43
Cont’d
4_____ What is the isoelectric point of arginine?
a) 5.60
b) 7.00
c) 7.90
d) 10.75
5._____ A solution with a pH of 2.2 contains 6.0 mmol of arginine. If 12.0 mmol of NaOH is added to the solution, what will be the pH after the NaOH has completely reacted with the arginine?
a) 14.00
b) 12.50
c) 10.75
d) 5.60

## Slide 44
You have titrated the a.a. -glutamic acid- with strong acid:
Draw a reasonable titration curve, label the axis and significant transitions
pKaα-COOH= 2.19 , pKaα-amino = 9.4 , and pKaR-COOH=4.25
b) Calculate the isoelectric point?
pI = (pKaα-COOH + pKaR-COOH) /2
pI = (2.19 + 4.25) /2 = 3.2

## Slide 45
Describe how you would prepare a 1 L of the following buffer 0.025M formic acid (mwt =178.93 g/mol) / sodium formate (mwt = 68.01 g/mol) buffer, pH = 4.0, containing 0.05M glucose (mwt= 180.156g/mol). (pKa =3.75)
pH = pKa + log [sodium formate (A-)]/ [formic acid(HA)]
Assume [A-] = y , [HA] = 0.025 –y
4= 3.75 + log y /( 0.025-y)------- 1.77=y / (0.025-y). (after antilog)
[A-] = y = 0.016 M
[HA] = 0.025 - 0.016 = 9 x 10-3 M
no. of moles of A- = M x V (L) = 0.016 x 1 = 0.016 moles
no. of moles of HA = M x V (L) = 9 x 10-3 x 1 = 9 x 10-3 moles
wt. of A- = no. of moles x mwt = 0.016 x 68.01 = 1.088 g
wt of HA= no. of moles x mwt = 9 x 10-3 x 178.93 = 1.610 g
Note: any additions of salts or sugars (that does not release H+or OH-) such as NaCl, EDTA, CaCl2 to buffers will not affect pH of solution.
Components of 0.025M formate
buffer, pH =4

## Slide 46
** This buffer contains glucose, its calculated weight is added to the buffer components where it not
effect the pH of buffer since it is a sugar that does not dissociate to H+ or OH- **
no. of moles of glucose = M x V (L) = 0.05 x 1= 0.05 moles
wt of glucose= no. of moles x mwt = 0.05 x 180.156 = 9.0078 g
To prepare 1L of this buffer ---------dissolve 1.088g of sodium formate, 1.61g of formic acid and 9.0078g of glucose in some water, then complete to the final volume of 1 liter and check pH.

## Slide 47: REVISION

## Slide 48
4.9g of CH3COOK (mwt= 98 g/mol) is dissolved in 125 cm3 of 1M CH3COOH and the solution was made up to 250ml, pKa=4.7.
a) Calculate the pH of the final solution.
pH = pKa + log [CH3COOK (A-)/ CH3COOH (HA)]
(no. of moles is calculated instead of concentration since it is a ratio)
no. of moles of A- = wt / mwt = 4.9 / 98 = 0.05 moles
no. of moles of HA= M x V (L) = 1 x 0.125 = 0.125 moles
( V= 125 cm3 =125 ml = 0.125L)
pH = 4.7 + log 0.05/0.125
pH = 4.3
b) Molarity of the buffer
M buffer = no. of moles of buffer / V of buffer (L)
No. of moles of buffer = no of moles A- + no. of moles HA
= 0.05 + 0.125 = 0.175 moles
M buffer = 0.175 / 0.25 = 0.7M

## Slide 49
Blood plasma at pH 7.4 contains 2.4 x 10-2 M H2CO3 and 1.2 x 10-3 M HCO3-. Calculate the pH after the addition of 3.2 x 10-3 M H+. Assume that the concentration of dissolved H2CO3 is maintained constant at 2.4 x 10-2 M (pKa =6.1)?
HCO3- + H+ ------------H2CO3
[HCO3-] = 1.2 x 10-3 + 3.2 x 10-3 = 4.4 x 10-3 M
pH = pKa + log (HCO3-/H2CO3)
pH = 6.1 + log (0.0044/0.024)
pH = 5.32
Design a shortcut method for preparing a 0.5M Phosphate buffer, pH = 7.0, where only one form of phosphate is provided?

## Slide 50
An enzyme –catalyzed reaction was carried out in a solution containing 0.2M Tris buffer. The pH of the reaction mixture at the start was 7.8. As a result of the reaction, 0.03mole/liter of H+ was produced. (pKa of Tris = 8.1)
a) What was the ratio of Tris+ (conjugate acid) / Tris0 (conjugate base) at the start of the reaction?
pH = pKa + log Tris0/ Tris+
7.8 = 8.1 + log Tris0/ Tris+
-0.3 =log Tris0/ Tris+
+0.3 =log Tris+/ Tris0 (to get red of –ve , multiply by (-1) and reverse the fraction in the opposite side)
+ 0.3 = log Tris+/ Tris0 (take antilog)
Tris+/ Tris0 = antilog of 0.3 = 2
The ratio is Tris+/ Tris0 =2/1

## Slide 51
b) What are the concentrations of Tris + and Tris0 at the start of the reaction?
-From the ratio above: Tris + = 2 , Tris 0 =1 -------- total components in the solution =2+1=3
-For [Tris+], multiply the concentration of buffer by 2/3
[Tris+] = 𝟐/𝟑 x 0.2 = 0.133M
For [Tris0], multiply the concentration of buffer by 1/3
[Tris0] = 𝟏/𝟑 x 0.2 = 0.067M
Check: The pH is less than the pKa ; [conjugate acid] > [conjugate base]; 0.133M>0.067M
c) Write the chemical equation showing how the Tris buffer maintained a near constant pH.
-The conjugate base reacts with the excess H+:
Tris0 + H+ ------------Tris+
What are the concentrations of Tris+ and Tris0 at the end of the reaction?
As a result of the reaction, 0.03 mole/ liter of H+ was produced, the amounts of Tris+ and Tris0
change as (Tris0 reacts with H+ released to produce more Tris+):
[Tris+] = 0.133 + 0.03 = 0.163 M
[Tris0] = 0.067 – 0.03 = 0.037 M

## Slide 52
e) What was the final pH of the reaction mixture?
pH = pKa + log Tris0/ Tris+
pH = 8.1 + log 0.037/0.163
pH= 7.456
f) What would the final pH be if no buffer were present?
If no buffer were present, the production of 0.03M H+ would bring the pH to:
pH = - log [H+]
pH = -log 0.03
pH = 1.52

## Slide 53
What volume of glacial acetic acid (density 1.06g/ml, mwt= 60.05g/mol) and what weight of solid potassium acetate (mwt= 98.142 g/mol) are required to prepare 5L of 0.2M acetate buffer, pH = 5.0, pKa =4.77?
pH = pKa + log CH3COOK (A-)/CH3COOH (HA);
Assume [A-] = y , [HA] = 0.2 –y
5 = 4.77 + log [y / (0.2 –y)]
0.23 = y / (0.2 –y). (after antilog) 0.3396 – 1.698 y = y
0.3396 = 2.698y
y = 0.3396/2.698
[A-] = y = 0.126M ; [HA] = 0.2 – 0.126= 0.074 M
no. of moles of A- = M x V (L) = 0.126 x 5 = 0.63 moles
no. of moles of HA = M x V (L) = 0.074 x 5 = 0.37 moles
wt. of A- = no. of moles x mwt = 0.63 x 98.142 = 61.8 g
wt of HA= no. of moles x mwt = 0.37 x 60.05 = 22.2 g
Note: To calculate volume of glacial acetic acid (HA): d = wt (g) / V(ml)
V(ml) = wt(g) / d
V of HA= 22.2 / 1.06 = 20.9 ml
To prepare 5L this buffer----------dissolve 61.8 g of potassium acetate in 20.9 ml of glacial acetic acid, then complete to the final volume of 5 liter and check pH.

## Slide 54
Describe the preparation of 40 L of 0.02M of phosphate buffer, pH 6.9 starting from solid Na3PO4 (mwt =163.9 g/mol) and 1M HCl.
Buffer components are: H2PO4- and HPO4-2, pKa =7.21 “the closest to the desired pH”
Na3PO4 + HCl ------- Na2HPO4
Na2HPO4 + HCl ------- Na2HPO4 (HPO4-2 ) + NaH2PO4 (H2PO4- )
Concentration is calculated by:
pH = pKa + log [(HPO4-2)/(H2PO4-)] ;
Assuming: [A-] = y , [HA] = 0.02 –y
[A-] = y = 6.57 x 10-3 M ; [HA] = 0.02 - 6.57 x 10-3 = 0.0134 M
no. of moles of A- = M x V (L) = 6.57 x 10-3 x 40 = 0.016 moles
no. of moles of HA = M x V (L) = 0.0134 x 40 = 0.536 moles
no. of moles of buffer = 0.016 + 0.536 = 0.7988 moles

## Slide 55
Start with 0.7988 moles of Na3PO4 + 0.7988 moles of HCl ------- 0.7988 moles of Na2HPO4
0.7988 moles of Na2HPO4 + 0.536 moles of HCl ------- 0.016 moles Na2HPO4 + 0.536 moles of NaH2PO4
Thus: no. of moles of HCl added = 0.7988 + 0.536 = 1.3348 moles
V of HCl = no. of moles / M = 1.3348/ 1 = 1.3348 L
Wt of Na3PO4 = no. of moles x mwt = 0.7988 x 163.9 = 130.9 g
To prepare 40L this buffer-------------dissolve 130.9 g of Na3PO4 in 1.3348 L of HCl, then complete to the final volume of 40 liter and check pH.

## Slide 56
Do you have questions?