Capillary electrophoresis

---
Principle of electrophoresis
Electrophoresis is the process of moving charged molecule in
solution by applying an electric field across the mixture.
Where molecules are allowed to migrate.
Separation modes according to:
charge
size

---
1- According to charge
When charged molecules are placed in an electric field, they migrate toward either the positive (anode)
or negative (cathode) pole according to their charge.

---
2-according to size
Mice run faster through the forest than elephant

---
Capillary Electrophoresis
known as capillary zone electrophoresis (CZE), can be used to separate ionic species by their charge.
In traditional electrophoresis, electrically charged analytes move in a conductive liquid medium under the influence of an electric field.

---
Instrumentation
Capillary tube
    - Varied length (normally 25-50 cm).
    - Small bore and thickness of the silica play a role.
Controllable high voltage power supply
Two electrode
Two buffer reservoirs
Detector

---
CE instrument

---
Electroosmotic flow
Electroosmotic flow is observed when an electric field is applied to a solution in a capillary that has fixed charges on its interior wall.
Charge is accumulated on the inner surface of a capillary when a buffer solution is placed inside the capillary.
In a fused-silica capillary, silanol (Si-OH) groups attached to the interior wall of the capillary are ionized to negatively charged silanoate (Si-O-) groups at pH values greater than three.
Attracted to the negatively charged silanoate groups, the positively charged cations of the buffer solution will form two inner layers of cations (called the diffuse double layer or the electrical double layer) on the capillary wall.

---
Electroosmotic flow
The first layer is referred to as the fixed layer because it is held tightly to the silanoate groups.
The outer layer, called the mobile layer, is farther from the silanoate groups.
The mobile cation layer is pulled in the direction of the negatively charged cathode when an electric field is applied.
Since these cations are solvated, the bulk buffer solution migrates with the mobile layer, causing the electroosmotic flow of the buffer solution.

---
Electroosmotic flow

---
Electroosmotic flow

---
pH, Silanol Population, and the rate of EOF flow
At very low pH, not many silanols are ionized and the EOF is slow.
As pH increases the number of ionized sites also increases. The EOF speed rises steadily.
At very high pH values, a maximum number of ionized sites is reached. The EOF speed also reaches a maximum

---
electrophoretic mobility and electro-osmotic flow
The separation of compounds by capillary electrophoresis is dependent on the differential migration of analytes in an applied electric field.
The electrophoretic mobility is proportional to the ionic charge of a sample.\
Electro-osmotic flow of the buffer solution is generally greater than that of the electrophoretic flow of the analytes.

---
all analytes are carried along with the buffer solution toward the cathode.
Negatively charged analytes are retained longer in the capillary due to their conflicting electrophoretic mobilities. small multiply charged cations migrate quickly and small multiply charged anions are retained strongly.

---
optimizing electrophoresis
Optimal electrophoretic separations must balance :
speed and resolution.
Higher voltage increases speed, but heat causes evaporation of the buffer and may denature molecules.
Higher ionic strength (buffer) increases conductivity, but enhances endosmotic effects.
[Speaker Notes: The conductivity (or specific conductance) of an electrolyte solution is a measure of its ability to conduct electricity.]

---
Joule Heating (heating due to voltage)
How to prevent Joule Heating?
use a smaller-bore capillary (less current reduces the heat produced)
a longer capillary (more surface area dissipates the heat generated).
The use of low conductivity buffers is also helpful in this regard.
effective cooling systems are required to ensure heat removal. Liquid cooling is the most effective means of heat removal and capillary temperature control.

---
Different between HPLC and CE

---
Different between HPLC and CE

---
application
Analysis of carbohydrates
Analysis of inorganic anions/metal ions
DNA profiling
Protein identification

---
Advantages &disadvantages of CE:
Advantages
    - Fast
    - Small Sample and solvent requirements
    - Relatively inexpensive
    - Automated
    - High efficiency and resolution
    - covers a broad range of applications in a wide variety of industries

---
Advantages &disadvantages of CE:
Disadvantages
    - Cannot identify neutral species
    - Joule Heating
    - Cannot discern shape