## Slide 1: Calibration of Micro-pipette By Gravimetric Method Nihal Gunasekara Sri Lanka
Bangladesh BEST Programme

## Slide 2: Types of Micro-pipettes
Single Channel
Multi- Channel
Bangladesh BEST Programme

## Slide 3: Requirements
Unit of volume Cubic millimeter : µl
Ref. temperature 20 °C
  To deliver Apparatus is calibrated to accurately transfer the stated volume to another container.
Bangladesh BEST Programme

## Slide 4: Apparatus Needed
Balance
Thermometer
Hygrometer
  Barometer
  Distilled or deionized water
Weighing vessel
Bangladesh BEST Programme

## Slide 5: Recommended Apparatus - Balance
Selected Volume (V) | Resolution of Balance (mg) | Std. Deviation (mg)
1 µl ≤ V ≤ 10 µl | 0.001 | 0.002
10 µl < V ≤ 100 µl | 0.01 0.02 | 
100 µl < V ≤ 1000 µl | 0.1 0.2 | 
1 ml < V ≤ 10 ml | 0.1 0.2 | 
Bangladesh BEST Programme

## Slide 6: Recommended Apparatus/Liquids
Apparatus/Liquid | Uncertainty
Thermometer | ≤ 0.2 °C
Hygrometer | ≤ 10%
Barometer | ≤ 0.5 kPa
Calibration Liquid | Distilled /de-ionized water
Bangladeshi BEST Programme

## Slide 7: Conditions for the Test
Temperature Maintained at 20 ± 0.5 °C and during calibration ± 0. 2 °C stability
Humidity Maintained > 50 %
Prefer to maintain at 70 ± 10 %
Stabilization Volumetric instruments and test water
1 to 2 hours
Evaporation Special precautions below 50 µl
Test cycle Less than 60 s
M
Bangladesh BEST Programme

## Slide 8: Test Volume
Fixed Volume Nominal Volume
Variable Volume At least three volumes
Preferably 10%, 50 % and nominal volume
Measurements 10 at each volume
M
Bangladesh BEST Programme

## Slide 9: Calibration Procedure- Single Channel
  Place the test liquid in the weighing vessel to a depth of at least up to 3mm and record the temperature, barometric pressure & humidity
  Fit the selected disposable tip to the micro pipette barrel
  Fill the tip with test liquid and expel to waste five time to reach a humidity equilibrium in the dead air volume
  Place the weighing vessel with its added water on the balance pan
Bangladesh BEST Programme

## Slide 10: Calibration Procedure- Single Channel
  Prior to any calibration test make sure that all instruments reached the ambient temperature in the laboratory
  During the calibration maximum temperature variation in the laboratory should not exceed ±0.2oC.
  Be sure that all the apparatus use for calibration and the distilled water are at room temperature.
Bangladesh BEST Programme

## Slide 11: Calibration Procedure
Replace the disposable tip of the piston pipette
Fill the piston pipette with test liquid, immersing its delivery orifice 2 to 3 mm below the surface of water
Release the operating button slowly and draw the pipette vertically against the side wall of container .
Expel the water to waste to pre-wet the tip and refill the piston pipette as above.
Record the mass m0 of the weighing vessel or tare the balance to zero
Deliver the contents of the pipette into the weighing vessel by touching pipette tip against the container wall just above liquid surface at an angle 30° to 45°
Record the mass m1 of the weighing vessel and determine the quantity delivered
Repeat the above cycle for ten measurements
Bangladesh BEST Programme

## Slide 12: Calibration Procedure- Multi-Channel
  Procedure is similar to single channel micro pipette
  During the calibration each channel should be considered as a single channel
Bangladesh BEST Programme

## Slide 13
Conversion of Mass into Volume
The Z correction factor is a function of test temperature and air pressure
Bangladesh BEST Programme

## Slide 14: Z Correction Factors for Distilled Water (Z values in micro-litres per milligram )
°C | 80 | 85 | 90 | 95 | 100 | 101.3 | 105
17 17.5 | 1.0021 1.0022 | 1.0021 1.0022 | 1.0022 1.0023 | 1.0022 1.0023 | 1.0023 1.0024 | 1.0023 1.0024 | 1.0023 1.0024
18.0 18.5 | 1.0022 1.0023 | 1.0023 1.0024 | 1.0023 1.0024 | 1.0024 1.0025 | 1.0025 1.0025 | 1.0025 1.0026 | 1.0025 1.0026
19.0 19.5 | 1.0024 1.0025 | 1.0025 1.0026 | 1.0025 1.0026 | 1.0026 1.0027 | 1.0026 1.0027 | 1.0027 1.0028 | 1.0027 1.0028
20.0 20.5 | 1.0026 1.0027 | 1.0027 1.0028 | 1.0027 1.0028 | 1.0028 1.0029 | 1.0028 1.0029 | 1.0029 1.0030 | 1.0029 1.0030
21.0 21.5 | 1.0028 1.0030 | 1.0029 1,0030 | 1.0029 1.0031 | 1.0030 1.0031 | 1.0031 1.0032 | 1.0031 1.0032 | 1.0031 1.0032
22.0 22.5 | 1.0031 1.0032 | 1.0031 1.0032 | 1.0032 1.0033 | 1.0032 1.0033 | 1.0033 1.0034 | 1.0033 1.0034 | 1.0033 1.0034
Temp. Air Pressure kPa
Bangladesh BEST Programme

## Slide 15
Mean Volume
Bangladesh BEST Programme

## Slide 16: Data Recording
Use spread sheet for data recording
Record Temperature, Humidity, Barometric pressure
Record final data in a data recording sheet
Review from a third party
Prepare the final calibration certificate
Bangladesh BEST Programme

## Slide 17
Thank you
Bangladesh BEST Programme